## Test Summary
Product QC Testing Summary
Product Catalog/Part No.: | XP3COV2FLURSV10 |  |  | (Select using drop-down menu.) |  |  | Type of Resample |  | N/A
Cartridge No. (LSP) | 12345 |  |  | Expiration Date |  | 2026 09 09
Table 1; Test and Control Material Information
Flu A Positive Control |  |  | Flu B Positive Control |  |  | RSV B Positive Control |  |  | SARS-CoV-2 Positive Control |  |  | Negative Control |  |  | Positive Control
Part No. | Lot No. | Exp Date | Part No. | Lot No. | Exp Date | Part No. | Lot No. | Exp Date | Part No. | Lot No. | Exp Date | Part No. | Lot No. | Exp Date | Part No. | Lot No. | Exp Date
525-00097 or 325-00070 | GJAN24AMasa | 2027-03-04 | 525-00099 or 325-00072 | 01JES34T | 2025 11 29 | 525-00102 or 325-00075 | 204Mate | 2026 01 01 | 025-00026 or 525-00169 | ABC021MoW | 2026-04-04 | 001-1471 or 500-2169 | 345235 | 2026 03 05 | 001-1471 | 345278 | 2026 09 09
Table 2; Pipette Information
Cepheid ID | Nr. 1 |  | Nr. 2 |  | Nr. 3 |  | N/A |  | N/A |  | N/A
Calibration Due Date | Jan-26 |  | Jan-26 |  | Mar-26 |  | N/A |  | N/A |  | N/A
Table 3; GeneXperts Used:
Cepheid ID | Infinity-VI |  | Infinity-VIII |  | GX1 |  | Infinity-I |  | N/A |  | N/A
Calibration Due Date | Jan-23 |  | Jan-23 |  | Feb-26 |  | Jan-23 |  | N/A |  | N/A
Cepheid ID | N/A |  | N/A |  | N/A |  | N/A |  | N/A |  | N/A
Calibration Due Date | N/A |  | N/A |  | N/A |  | N/A |  | N/A |  | N/A
Cepheid ID | N/A |  | N/A |  | N/A |  | N/A |  | Bag Numbers Tested Using Infinity |  | #01-10 (x209)
Calibration Due Date | N/A |  | N/A |  | N/A |  | N/A
Table 5; Single Minor Failure Mode: | [7 empty] | Table 6; Visual and Delamination Resample Results
Type |  | Original Count | Result | Resample Count | Final Count | Result |  | Type |  | Resample Count
Invalid (w/o error code) and SPC Ct=0 |  | 0 | Pass | N/A | N/A | N/A |  | Minor Visual Failure |  | N/A
Curve Fit Error |  | 0 | Pass | N/A | N/A | N/A |  | Major Visual Failure |  | N/A
Probe Check Low |  | 0 | Pass | N/A | N/A | N/A |  | Delamination |  | N/A
Probe Check High |  | 0 | Pass | N/A | N/A | N/A |  | Note: Enter in results if applicable (all resamples other than functional and full)
CIT and SLD |  | 0 | Pass | N/A | N/A | N/A
Pressure Abort or Max Pressure > 90 PSI |  | 0 | Pass | N/A | N/A | N/A |  | Table 7; Major Failure Results
Sample Volume Adequacy |  | 0 | Pass | N/A | N/A | N/A |  | Type |  |  | Original Count | Resample Count
Delamination |  | 0 | Pass | N/A | N/A | N/A |  | No. False Negative |  |  | 0 | N/A
[8 empty] | No. False Positive |  |  | 0 | N/A
[8 empty] | No. Major Visual Inspection Failures: |  |  | 0 | N/A
| Max Count | 0 |  |  | N/A
| Max Acceptable: | 4 |  |  | 10
| (Out of ) | (210) |  |  | (410)
Table 8; Result Summary:
Type |  |  | Original Count | Accept Criteria | Result |  | Resample Count | Accept Criteria | Combined Count | Result
Total No. Cartridges Registered: |  |  | 0 | N/A | N/A |  | N/A | N/A | N/A | N/A
Total No. Valid Test Replicates: |  |  | 0 | N/A | N/A |  | N/A | N/A | N/A | N/A
No. Minor Failures: |  |  | 0 | N/A | N/A |  | N/A | N/A | N/A | N/A
No. Seal Test Failures: |  |  |  | ≤9 | Fail |  | N/A | N/A | N/A | N/A
No. Total Major Failures: |  |  | 0 | ≤1 Resample at 2-3 | Pass |  | N/A | Combined ≤5 | N/A | N/A
No. Total Failures: |  |  | 0 | ≤9 | Pass |  | N/A | Resample Data Only ≤10 | N/A | N/A
|  | Original Result: | Fail |  |  | NC# | N/A
|  | Overall Result | Fail |  |  | LFI# | N/A
Comments
N/A
| Recorded By: | Jesper Fredriksson | [5 empty] | Date: | 2025 11 20
| PQC Reviewed By: | Namn Namnsson | [5 empty] | Date: | 45982