## Test Summary
Product QC Testing Summary
Product Catalog/Part No.: | GXMTBRIFULTRA50 |  |  | (Select using drop-down menu.) |  |  | Type of Resample |  | NA
Cartridge No. (LSP) | 78808 |  |  | Expiration Date |  | 2026 12 27
Table 1; Test and Control Material Information
TB1081-MDR2 Cells |  | TET Buffer |  | Sample Reagent |  |  | Sample Reagent
Part No.* | Lot No. | Part No.* | Lot No. | Part No.* | Lot No. | Bag # | Part No.* | Lot No. | Bag #
001-1923 | C15NOV23A | 500-1315 or 500-0995 | 1001487581 | 500-2095, 500-2290, 500-2608, 700-5208, 700-6052 or 700-9287,700-8870 | 1001490276 (C252400039) | NEG: 00-03 POS: 01-04 | 500-2095, 500-2290, 500-2608, 700-5208, 700-6052 or 700-9287,700-8870 | 1001488414 | NEG: 04-10 POS: 05-10 WT
TBWT Cells |  |  |  | Sample Reagent |  |  | Sample Reagent
Part No.* | Lot No. |  |  | Part No.* | Lot No. | Bag # | Part No.* | Lot No. | Bag #
001-1968 | B01NOV24B |  |  | 500-2095, 500-2290, 500-2608, 700-5208, 700-6052 or 700-9287,700-8870 | N/A | N/A | 500-2095, 500-2290, 500-2608, 700-5208, 700-6052 or 700-9287,700-8870 | N/A | N/A
* Circle one, if more than 1 option is displayed.
Table 2; Pipette Information
Cepheid ID | Nr. 20 |  | Nr. 37 |  | Nr. 18 |  | Nr. 19 |  | Nr. 31 |  | N/A
Calibration Due Date | Mar-26 |  | Mar-26 |  | Mar-26 |  | Mar-26 |  | Mar-26 |  | N/A
Table 3; GeneXperts Used:
Cepheid ID | Infinity-V |  | Infinity-I |  | Infinity-VI |  | Infinity-VIII |  | N/A |  | N/A |  | Bag #s Tested on Inifinity | N/A
Calibration Due Date | Mar-26 |  | Mar-26 |  | Mar-26 |  | Mar-26 |  | N/A |  | N/A
Cepheid ID | N/A |  | N/A |  | N/A |  | N/A |  | N/A |  | N/A
Calibration Due Date | N/A |  | N/A |  | N/A |  | N/A |  | N/A |  | N/A
Table 4: Single Minor Failure Mode: | [7 empty] | Table 5: Visual and Delamination Resample Results
Type |  | Original Count | Result | Resample Count | Final Count | Result |  | Type |  | Resample Count
Invalid (w/o error code) and SPC Ct=0 |  | 0 | Pass | NA | NA | NA |  | Minor Visual Failure |  | N/A
Curve Fit Error |  | 0 | Pass | NA | NA | NA |  | Major Visual Failure |  | N/A
Probe Check Low |  | 2 | Pass | NA | NA | NA |  | Delamination |  | N/A
Probe Check High |  | 0 | Pass | NA | NA | NA |  | Note: Enter in results if applicable (all resamples other than functional and full)
CIT and SLD |  | 1 | Pass | NA | NA | NA
Pressure Abort or Max Pressure > 90 PSI |  | 0 | Pass | NA | NA | NA |  | Table 6: Major Failure Results
Sample Volume Adequacy |  | 0 | Pass | NA | NA | NA |  | Type |  |  | Original Count | Resample Count
Delamination |  | 0 | Pass | NA | NA | NA |  | No. False Negative |  |  | 0 | NA
[8 empty] | No. False Positive |  |  | 0 | NA
[8 empty] | No. Major Visual Inspection Failures: |  |  | 0 | NA
| Max Count | 2 |  |  | NA |  |  | MTB Trace Detected |  |  | 0 | NA
| Max Acceptable: | 4 |  |  | 10
| (Out of ) | (210) |  |  | (410)
Table 7: Result Summary:
Type |  |  | Original Count | Accept Criteria | Result |  | Resample Count | Accept Criteria | Combined Count | Result
Total No. Cartridges Registered: |  |  | 210 | NA | NA |  | NA | NA | NA | NA
Total No. Valid Test Replicates: |  |  | 210 | NA | NA |  | NA | NA | NA | NA
No. Minor Failures: |  |  | 8 | NA | NA |  | NA | NA | NA | NA
No. Seal Test Failures: |  |  | 0 | ≤9 | Pass |  | NA | NA | NA | NA
No. Total Major Failures: |  |  | 0 | ≤ 1 Resample at 2-3 | Pass |  | NA | ≤ 5 | NA | NA
No. Total Failures: |  |  | 8 | ≤ 9 | Pass |  | NA | ≤ 10 | NA | NA
|  | Original Result: | Pass |  |  | NC# | N/A
|  | Overall Result: | Pass |  |  | LFI# | N/A
| Comments | [9 empty] | ¨N/A
Loten testad av: Rafail Ramolis, Noor Nawaf och Bincenth Salazar
| Recorded By: | Alen Devetak | [5 empty] | Date: | 2025 07 22
| PQC Reviewed By: | Maria Bouikou | [5 empty] | Date: | 2025 07 23